AT30 series ASCII protocol and memory map

Note1 – ACSII protocol commands and hex values are not case sensitive.

Note2 – terminal program used for this document:

Note3 – Terminal program using scripting to send ACSII protocol commands and hex values as a string.

AT30 series ASCII protocol Debug Commands index:

; AD - ADDRESS DEBUG

;  AD,HH WILL ENABLE DEBUG IF HH = SYSTEM I.D. IN EEPROM, DISABLE IF UNEQUAL

; AP - DUMP ARCHIVE POINTER

;  AP   WILL RESPOND 4 HEX - NUMBER OF ENTRIES IN ARCHIVES

; AR - DUMP ARCHIVES

;  AR,HHHH   WILL DUMP 4 ARCHIVES (32 CHARACTERS) AS STRING OF HEX

;     STARTS AT THE HHHH 'TH ARCHIVE

; DI - DUMP DATA FROM DISPLAY

;  DI    WILL SET DISFL,  WAIT TILL CLEARED , THEN DO A DD OF DISBUF

; KP - ACT LIKE KEYBOARD

;  KP,(CHARACTERS 0-9, *, #)  FAKES AS IF THOSE CHARACTERS ENTERED THRU KEYPAD

;    ANY CHARACTER OTHER THAN 0 THRU 9, # AND * WILL TERMINATE the KP command

DR,XXXX,XXXX (retrieve ram and EEProm data from shared RAM)

XXXX,XXXX (Start and end memory location)

Memory location in shared RAM for TDR RAM (0000-04FF)

Memory location in shared RAM for TDR EEProm (2000,21FF)

Returns: <MMMM> <D1> <D2> … <DN> (where MMM is starting memory location and D? are corresponding data bytes up to end location.

DP,Cable#,Map#,XXXX,XXXX (retrieve Cable Map data, Current Map data, Temperature Shift Map data and cable setup and flag data)

Cable# 1-4 (Cable 2 through 4 only apply to AT30K panels).

Map# 1-7

XXXX,XXXX (Start and end memory location)

Memory location for cable Map data (6218-69d7)

Memory location for Current map (4040-4fff), Temp Shift Map (5040-5fff) and cable setup and flag data (6000-61f7)

; PW,XXXXXXXX PASSWORD ENTRY AND CHECK (3179636424)

; CS - Dump Cable Status

; PV - Dump TDR Firmware Version

; CV - Dump Comm Firmware Version

; UV - Dump USB Firmware Version

To edit TDR RAM and TDR EEProm data use the JR command:

JR,XXXX,XX - JAM MEMORY RAM and EEPROM(ENTER DATA WITHOUT CHECKING IT)

XXXX (Memory location to edit)

xx (hex byte or spacebar = next memory location or , = previous memory location or .  = end edit)

use TDR address 0400-04FF for EEProm, TDR address 1000-14FF Ram

; FR,XXXX,XX - Fill MEMORY RAM and EEPROM(ENTER DATA WITHOUT CHECKING IT) use main address 0400-04FF EEProm,1000-14FF Ram

;                 enter with start and end address already stored, just get value

To edit cable/senor setup and flag data use the JF command:

JF,Cable#,XXXX,xx

Cable# 1-4 (Cable 2 through 4 only apply to AT30K panels).

XXXX (Memory location to edit)

xx (hex byte or spacebar = next memory location or , = previous memory location or .  = end edit)

Can only edit cable/sensor setup and flag data (6000-61ff)

When the editing process is completed enter the WD (Watchdog) command.

WD - Watchdog- resets com and tdr (after the WD command is sent communication to the panel must be reestablished).

AT30 series ASCII protocol Debug Commands detail with examples and detailed memory map:

To clear ACSII protocol command to new prompt, enter Esc key or 1B (hex code for the Esc key)

Terminal response if communication port is open:

11/7/2018 10:35:13.812 [TX] - <ESC>_

11/7/2018 10:35:13.842 [RX] - <CR><LF>

MON 30A>

Terminal response if communication port is closed (listen only mode):

11/8/2018 09:06:56.890 [TX] - <ESC>_

To open communication to the AT30 series panel open a terminal screen to the selected panel via Ethernet or serial port (RS-485 network) enter - Esc key or 1B (hex code for the Esc key) then enter – AD,(system I.D. number in hex)

Example for addressing system 1, enter -  AD,01

Terminal response:

11/9/2018 13:52:30.187 [TX] - ad,01

11/9/2018 13:52:30.241 [RX] - <CR><LF>

MON 30A>

To enable restricted access, enter PW,3179636424

Terminal response:

11/8/2018 11:50:58.459 [TX] - pw,3179636424

11/8/2018 11:50:58.503 [RX] - pw,XXXXXXXXXX ENABLED<CR><LF>

MON 30A>

To close communication (listen only mode) to the AT30 series panel, enter – AD,FF

Terminal response:

MON 30A>

11/8/2018 11:52:00.371 [TX] - ad,ff

11/8/2018 11:52:00.416 [RX] - ad,ff

EEPROM Memory Map:

Note 1 – EEPROM data from the TDR uC is viewed from a copy in shared RAM (address 2000 – 21ff).

Note 2 – when EEPROM data is edited it must be directly written to the EEPROM address in the TDR uC (address 0400 – 05ff).

PASSWORD TABLE (Shared RAM address 2000 to 20e7, TDR uC EEPROM address 0400 to 04e7) - 29 ENTRIES OF 8 BYTES EACH

FIRST 5 BYTES ARE HEX OF 10 DIGIT ENTRY

- EACH STORED AS $E-VALUE (E=0, D=1, C=2, ETC UP TO 5=9; NO ENTRY STAYS $F)

- Example 7089662190 at 250 level = 7E 65 88 CD 5E FF FF FA

- NEXT 2 BYTES ARE SPARE FOR NOW

- LAST BYTE IS LEVEL OF AUTHORIZATION PERMITTED

To display password table, enter - DR,2000,20e7

Terminal response:

11/9/2018 13:56:04.840 [TX] - dr,2000,20e7

11/9/2018 13:56:04.897 [RX] - dr,2000,20e7<CR><LF>

2000  7E 65 88 CD  5E FF FF FA  FF FF FF FF  FF FF FF FF  ~e.M^..z........<CR><LF>

2010  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2020  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2030  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2040  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2050  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2060  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2070  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2080  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2090  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20A0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20B0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20C0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20D0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20E0  FF FF FF FF  FF FF FF FF  ........<CR><LF>

MON 30A>

To edit password table:

Example – add password 111111 at level 250 to second password memory locations,

enter jr,0408dd dd dd     fa.

Response to terminal:

MON 30A>

11/9/2018 14:19:31.042 [TX] - jr,0408dd dd dd     fa.

11/9/2018 14:19:31.071 [RX] - jr,0408<CR><LF>

0408 dd  <CR><LF>

0409 dd  <CR><LF>

040A dd  <CR><LF>

040B  <CR><LF>

040C  <CR><LF>

040D  <CR><LF>

040E  <CR><LF>

040F fa .<CR><LF>

MON 30

To display edited password table, enter - DR,2000,20e7

Terminal response:

MON 30A>

11/9/2018 14:17:49.855 [TX] - dr,2000,20e7

11/9/2018 14:17:49.889 [RX] - dr,2000,20e7<CR><LF>

2000  7E 65 88 CD  5E FF FF FA  DD DD DD FF  FF FF FF FA  ~e.M^..z]]]....z<CR><LF>

2010  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2020  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2030  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2040  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2050  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2060  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2070  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2080  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

2090  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20A0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20B0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20C0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20D0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

20E0  FF FF FF FF  FF FF FF FF  ........<CR><LF>

MON 30A>

Demo Flag (Shared RAM address 20e8, TDR uC EEPROM address 04e8) –

- Always an FF for AT30 series systems.

- Demo mode is enable via the J4 header in the AT30 series panels (shunt horizontal across position 6)

Language Flag (Shared RAM address 20e9, TDR uC EEPROM address 04e9) –

- 3A=ENGLISH FT,3B=GERMAN,3C=ENGLISH M,3D=SPANISH

To display Language Flag, enter - DR,20e9,20e9

Terminal response:

MON 30A>

11/9/2018 14:54:03.414 [TX] - dr,20e9,20e9

11/9/2018 14:54:03.475 [RX] - dr,20e9,20e9<CR><LF>

20E9  3A  :<CR><LF>

MON 30A

To edit Language Flag, enter - JR,04e9,( 3A=ENGLIST FT,3B=GERMAN,3C=ENGLISH M, or 3D=SPANISH)(.)

Example for change language to Spanish, enter jr,04803d.

Response to terminal:

MON 30A>

11/9/2018 14:55:06.049 [TX] - jr,04e93d.

11/9/2018 14:55:06.119 [RX] - jr,04e9<CR><LF>

04E9 3d .<CR><LF>

MON 30A>

To display edited Language Flag, enter - DR,20e9,20e9

Terminal response:

MON 30A>

11/9/2018 14:56:29.850 [TX] - dr,20e9,20e9

11/9/2018 14:56:29.877 [RX] - dr,20e9,20e9<CR><LF>

20E9  3D  =<CR><LF>

MON 30A>

Autosilence Flag (Shared RAM address 20EA, TDR uC EEPROM address 04EA) – 55 disable

- Example

  - Get Autosilence Flag: DR,20EA,20EA

    - Return: 20EA NN

      - NN: 55 == Mentoring Disabled , !55 == Monitoring Enabled

  - Disable Autosilence: JR,04EA<>55.

  - Enable Autosilence: JR,04F0<>FF.

Alarm Silence Time (Shared RAM address 20EC-20EDputt, TDR uC EEPROM address 04EC-04ED) – 55 disable

  - Get Alarm Silence Time: DR,20EC,20ED

    - Return: 20EA N1 N2

      - N1 = Hours in Minutes in hex (up to 96 (0x60) allowed) N2 = Minutes in hex

  - Set silence time to 3.5 hours: JR,04EC<>03 <>1e. (note <> indicates panel will respond with address)

Monitor Flag (Shared RAM address 20f0, TDR uC EEPROM address 04f0)

- AA – Monitoring enabled, anything else – Monitoring disable.

- Example

  - Monitor Flag: DR,20F0,20F0

    - Return: 20F0 NN

      - NN: AA == Mentoring Enabled,! AA == Monitoring Disabled

  - Disable Monitoring: JR,04F0<>01.

  - Enable Monitoring: JR,04F0<>AA.

Serial Port 1 Data Type (Shared RAM address 20f1, TDR uC EEPROM address 04f1)

- 01=ASCII, 02=ModBus, 03=BacNet (not active)

- Example

  - Get Serial Port 1 Data Type: DR,20F1,20F1

    - Return: 20F1 NN

    - NN:  01 == ASCII, 02 == MODBUS

  - Set Serial Port 1 Data Type to ASCII: JR,04F1<>01.

  - Set Serial Port 1 Data Type to MODBUS: JR,04f1<>02. (note <> indicates panel will respond with address)

Serial Port 2 Data Type (Shared RAM address 20f2, TDR uC EEPROM address 04f2)

- Example

  - Get Serial Port 2 Data Type: DR,20F2,20F2

    - Return: 20f2 NN

      - NN: 01=ASCII, 02=ModBus, 03=BacNet (not active)

  - Set Serial Port 2 Data Type to ASCII: JR,04F201.

  - Set Serial Port 2 Data Type to MODBUS: JR,04f2<>02. (note <> indicates panel will respond with address)

Serial Port 1 Baud Rate (Shared RAM address 20f8, TDR uC EEPROM address 04f8)

- Example

- Get Serial Port 1 Baud: dr,20f8,20f8

  - Returns: 20f8 NN

    - NN: 00=1200, 01=2400I, 02=4800, 03=9600, 04=19200, 05=38400, 06=57600 and 07=115200

- Set Serial Port 1 Baud to 57600: jr,04F8<>06. (note <> indicates panel will respond with address)

Serial Port 2 Baud Rate (Shared RAM address 20f9, TDR uC EEPROM address 04f9)

- Example

  - Get Serial Port 2 Baud: dr,20f9,20f9

    - Returns: 20f9 NN

      - NN: 00=1200, 01=2400I, 02=4800, 03=9600, 04=19200, 05=38400, 06=57600 and 07=115200

  - Set Serial Port 2 Baud to 57600: jr,04F9<>06. (note <> indicates panel will respond with address)

Relay Normal State Flag (Shared RAM address 20fc, TDR uC EEPROM address 04fc)

- 10(hex) = cable relays normal state – energized, common relay normal state – not energized

- 04(hex) = cable relays normal state – not energized, common relay normal state – energized

- 14(hex) = cable relays normal state – energized, common relay normal state - energized

- 00(hex) = cable relays normal state – not energized, common relay normal state – not energized (default state)

- Example

- Get Relay Normal State: dr,20fC,20fC

  - Returns: 20fc NN

    - NN:

      - 10(hex) = cable relays normal state – energized, common relay normal state – not energized

      - 20fc 04(hex) = cable relays normal state – not energized, common relay normal state – energized

      - 20fc 14(hex) = cable relays normal state – energized, common relay normal state - energized

      - 20fc 00(hex) = cable relays normal state – not energized, common relay normal state – not energized (default state)

- Set Relay Normal State to Normal Active: jr,04FC<>14. (note <> indicates panel will respond with address)

System I.D. (Shared RAM address 20fd, TDR uC EEPROM address 04fd)

- Valid I.D. = 01 to fe (hex),  00 or ff (hex) com ports disabled

- Example:

  - Get ID: dr,20fD,20fD

    - Returns: 20FD NN (nn == Device ID in hex)

  - Set ID to 128 dec: jr,04FD<>80. (note <> indicates panel will respond with address)

IP Address (Shared RAM address 2100 - 2103, TDR uC EEPROM address 0500 - 0503)

- 4 hex bytes FF.FF.FF.FF = 255.255.255.255

  - Get IP: dr,2100,2103

    - Returns: 2100 N1 N2 N3 N4 (N1 – N4 is the IP address from left to right 192.168.0.1 is C0 A8 00 01)

  - Set IP to 1.2.3.4: jr,0500<>01 <>02 <>03 <>04. (Data can be loaded in reverse with “,” separator instead of” “ shown. note <> indicates panel will respond with address.)

Subnet Mask (Shared RAM address 2104- 2107, TDR uC EEPROM address 0504 - 0507)

- 4 hex bytes FF.FF.FF.FF = 255.255.255.255

  - Get Subnet: dr,2104,2107

    - Returns: 2104 N1 N2 N3 N4 (N1 – N4 is the IP address from left to right 255.255.255.0 is FF FF FF 00)

  - Set Subnet to 255.255.0.0: jr,0504<>FF <>FF <>00 <>00. (Data can be loaded in reverse with “,” separator instead of” “ shown. note <> indicates panel will respond with address.)

Gateway (Shared RAM address 2108 – 210b, TDR uC EEPROM address 0508 – 050b)

- 4 hex bytes FF.FF.FF.FF = 255.255.255.255

  - Get Gateway: dr,2108,210B

    - Returns: 2108 N1 N2 N3 N4 (N1 – N4 is the IP address from left to right 192.168.0.2 is C0 A8 00 01)

  - Set Gateway to 192.168.1.3: jr,0508<>C0 <>A8 <>00 <>03. (Data can be loaded in reverse with “,” separator instead of” “ shown. note <> indicates panel will respond with address.)

.

Auto-mapping Flag  (Shared RAM address 210d, TDR uC EEPROM address 050d)

  - Get Automap: dr,210D,210D

    - Returns: 210D NN (NN – AA == Auto Mapping Disabled !AA == Auto Mapping Enabled)

  - Set Automap Off: jr,050D<>AA. (note <> indicates panel will respond with address.)

MAC Address bytes 4 through 6 (Shared RAM address 2120 - 2122, TDR uC EEPROM address 0520 - 0522)

- 4th, 5th and 6th MAC address bytes (hex)

- Mac address generated by PAL-AT firmware (derived from panel serial number)

- 1st, 2nd and 3rd MAC address are hard coded to 10

- Example

  - Get Mac last 3 (first 3 hard coded 10-5c-3b): dr,2120,2122

    - Returns: 2020 N1 N2 N3 (N1-N3 MAC bytes 4-6)

  - Set MAC last 4 : jr,0520<>N1 <>N2 <>N3. (N1-N3 desired bytes 4-6 in hex. (note <> indicates panel will respond with address.))

Cable Relay Behavior (Shared RAM address 2137, TDR uC EEPROM address 0537)

- Cable relay changes state on leak only = aa (hex)

- Cable relay changes state on all faults = any hex number between 00 and ff except aa (default state)

- Example:

  - Get Cable Relay Behavior: dr,2137,2137

    - Return: 2137 NN (aa == active on leak only, !aa==active on any fault)

  - Set Cable Relay Behavior to leak only: jr,0537<>AA. (note <> indicates panel will respond with address.)

USE ABOVE EXAMPLES FOR DISPLAYING AND EDITING

System Serial Number (Shared RAM address 2139 -2141, TDR uC EEPROM address 0539-0541)

- 9 hex coded ACSII digits

- Example

  - Get Serial Number: dr,2139,2141

    - Return: 2139 N1 N2 N3 N4 N5 N6 N7 N8 N9(N1-N9 is serial number example for AT30K == 33 32 31 39 31 33 33 30 35 == “321913305”

  - Set Serial Number for AT30C 301913305: jr,0539<>33 <>30 <>31 <>39 <>21 <>33 <>33 <>30 <>35. (note <> indicates panel will respond with address.)

Serial number in following example - 321408801

To display serial number, enter dr,2139,2141

MON 30A>

11/16/2018 15:41:01.278 [TX] - dr,2139,2141

11/16/2018 15:41:01.304 [RX] - dr,2139,2141<CR><LF>

2139  33 32 31 34  30 38 38 30  31  321408801<CR><LF>

MON 30A>

Enable User Defined Splash Screen (Shared RAM address 2150 -2151, TDR uC EEPROM address 0550-0551)

- Enable User Defined Splash Screen = aa 55 (hex)

- Enable Default Splash Screen = any hex number between 00 00 and ff ff except aa 55

USE ABOVE EXAMPLES FOR DISPLAYING AND EDITING

User defined Splash Screen (Shared RAM address 2152 -217a, TDR uC EEPROM address 0552-057a)

- 34 ACSII user defined characters (at the end of 34 user defined character add 20 56 80 81 82 83 04 (hex)).

- If the user defined splash screen is less than 34 characters add space characters (20 hex) until there are 34 characters.

- Following is an example of adding ‘Test Splash Screen’ message to the PAL-AT

enter jr,055254 65 73 74 20 53 70 6c 61 73 68 20 53 60 72 65 65 6e 20 20 20 20 20 20 20 20 20 20 20 20 20 20 20 20 20 56 80 81 82 83 04.

Response to terminal:

MON 30A>

11/30/2018 15:11:22.967 [TX] - jr,055254 65 73 74 20 53 70 6c 61 73 68 20 53 63 72 65 65 6e 20 20 20 20 20 20 20 20 20 20 20 20 20 20 20 20 20 56 80 81 82 83 04.

11/30/2018 15:11:23.035 [RX] - jr,0552<CR><LF>

0552 54  <CR><LF>

0553 65  <CR><LF>

0554 73  <CR><LF>

0555 74  <CR><LF>

0556 20  <CR><LF>

0557 53  <CR><LF>

0558 70  <CR><LF>

0559 6c  <CR><LF>

055A 61  <CR><LF>

055B 73  <CR><LF>

055C 68  <CR><LF>

055D 20  <CR><LF>

055E 53  <CR><LF>

055F 63  <CR><LF>

0560 72  <CR><LF>

0561 65  <CR><LF>

0562 65  <CR><LF>

0563 6e  <CR><LF>

0564 20  <CR><LF>

0565 20  <CR><LF>

0566 20  <CR><LF>

0567 20  <CR><LF>

To display User Defined Splash Screen and the splash screen memory location, enter - DR,2150,217a

Terminal response:

MON 30A>

11/30/2018 15:11:56.671 [TX] - dr,2150,217a

11/30/2018 15:11:56.715 [RX] - dr,2150,217a<CR><LF>

2150  AA 55 54 65  73 74 20 53  70 6C 61 73  68 20 53 63  *UTest Splash Sc<CR><LF>

2160  72 65 65 6E  20 20 20 20  20 20 20 20  20 20 20 20  reen            <CR><LF>

2170  20 20 20 20  20 56 80 81  82 83 04       V.....<CR><LF>

MON 30A>

Queue Table (Shared RAM address 21E0 -21E, TDR uC EEPROM address 05E0-05EC)

- 21E0 – Cable(s) with queued events – 00 = no cables, 01 = cable 1, 02 = cable 2, 03 = cable 1 and 2, 04 = cable 3 … 0E = Cable 2, 3 and 4, 0F = cable 1,2 3, and 4

- 21E1 – number of queued events for cable 1

- 21E2 and 21E3 – where in the current map’s memory for cable 1 the queued event occurred (mask off 80’s bit from high order byte)

- 21E4 – number of queued events for cable 2

- 21E5 and 21E6 – where in the current map’s memory for cable 2 the queued event occurred (mask off 80’s bit from high order byte)

- 21E7 – number of queued events for cable 3

- 21E8 and 21E9 – were in the current map’s memory for cable 3 the queued event occurred (mask off 80’s bit from high order byte)

- 21EA – number of queued events for cable 4

- 21EB and 21EC – where in the current map’s memory for cable 4 the queued event occurred (mask off 80’s bit from high order byte)

To display Queue Table, enter - DR,21e0,21ec

Terminal response: (shows that only cable 3 has events in the queue, cable 1 and 2 never had a queue event, cable 3 has 2 queued events and that the last event occurred at memory location 954 of the current map for cable 3, cable 4 had a queued event that was cleared at memory location B2B of the current map for cable 4.)

MON 30A>

11/29/2018 14:23:19.481 [TX] - dr,21e0,21ec

11/29/2018 14:23:19.511 [RX] - dr,21e0,21ec<CR><LF>

21E0  04 00 00 00  00 00 00 02  19 54 00 1B  2B  .........T..+<CR><LF>

MON 30A>

To clear the queue, the user must have proper clearance and when the queue is cleared it must be logged with time, data and user name.

To clear the queue via Palcom Server, enter - jr,05e000 00   00   00   00.

MON 30A>

11/30/2018 07:52:40.626 [TX] - jr,05e000 00   00   00   00.

11/30/2018 07:52:40.727 [RX] - jr,05e0<CR><LF>

05E0 00  <CR><LF>

05E1 00  <CR><LF>

05E2  <CR><LF>

05E3  <CR><LF>

05E4 00  <CR><LF>

05E5  <CR><LF>

05E6  <CR><LF>

05E7 00  <CR><LF>

05E8  <CR><LF>

05E9  <CR><LF>

05EA 00 .<CR><LF>

MON 30A>

To display cleared Queue Table, enter - DR,21e0,21ec (note: last current map’s memory location does not clear, only the queue counters clear)

MON 30A>

11/30/2018 07:53:10.996 [TX] - dr,21e0,21ec

11/30/2018 07:53:11.025 [RX] - dr,21e0,21ec<CR><LF>

21E0  00 00 00 00  00 00 00 00  19 54 00 1B  2B  .........T..+<CR><LF>

MON 30A>

End of EEPROM Memory Map

Cable/Sensor Memory Map:

To display cable/sensor data use the DP command:

DP,X,X,XXXX,XXXX     DP,Cable#,Map#,XXXX,XXXX

Cable# 1-4 (Cable 2 through 4 only apply to AT30K panels).

Map# 1-7 (Reference Maps)

XXXX,XXXX (Start and end memory location)

Memory location for stored cable Reference Map data (6218-69d7)

Memory location for Current map (4040-4fff), Temp Shift Map (5040-5fff) and cable setup and flag data (6000-61f7)

To edit cable/senor setup data use the JF command:

JF,Cable#,XXXX,xx

Cable# 1-4 (Cable 2 through 4 only apply to AT30K panels).

XXXX (Memory location to edit)

xx (hex byte or spacebar – next memory location or , - previous memory location or .  – end edit)

Can only edit cable/sensor setup and flag data (6000-61ff)

When the editing process is completed enter JF,W then WD (Watchdog) command.

; WD - Watchdog- resets com and tdr

| Cable Type | Speed – feet/count | Speed – meters/count |
| --- | --- | --- |
| 1 – APT | 4.69 | 1.52 |
| 2 – AGW | 5.78 | 1.76 |
| 3 – JMP, JMP-U and JMP-UD | 4.59 | 1.41 |
| 4 – FGH | 5.93 | 1.81 |
| 5 – TFH | 5.77 | 1.76 |
| 6 – not used |  |  |
| 7- AGT | 4.97 | 1.52 |
| 8 – JPL | 4.59 | 1.41 |
| 9 – JPP | 5.72 | 1.75 |
| 10 – Probe | 4.59 | 1.41 |
| 11 – not used |  |  |
| 12 - not used |  |  |
| Cable Type | Speed – feet/count | Speed – meters/count |
| 13 – AGW Gold | 5.73 | 1.81 |
| 14- AGT Gold | 4.97 | 1.52 |
| 15 – TFH Gold | 5.77 | 1.76 |

Cable x Setup and Flag Data (x= 1 to 4): From address 6000 to 61f7:

6000 – Start pulse width – must always be 02 Hex.

6001 to 6002 – minimum data point that is recorded of any map - must always be 00 08 Hex. (include for Reference, Current and Temperature Shift Map data)

6003 to 6004 – Last data point that is recorded – 07 C0 Hex maximum (include for Reference, Current and Temperature Shift Map data)

6005 – must always be FF Hex.

6006 – Sensitivity – Low = 31 Hex, Medium = 11 Hex, High = 21 Hex (included in Cable/Sensor information on Palcom Server)

6007 – Resting Value – typically between 1e and 21 Hex.

6008 – must always be a 3F Hex

6009 to 600a – Monitor cable on/off – On = aa 55 Hex, Off = 55 aa Hex. (included in Cable/Sensor information on Palcom Server)

600b – Criteria Table – valid values – 03 Hex (Monitoring Hydrocarbon/ZBA installed/TFH installed/TFH Gold installed), 04 Hex (probe installed), 06 Hex – (ATP Cable installed) and 05 Hex – (all others) (included in Cable/Sensor information on Palcom Server)

600c – Which Map used – Map1 = 00 Hex, Map 2 = 01 Hex, … and Map 7 = 06 Hex.  Any hex number from 07 to ff = no map (included in Cable/Sensor information on Palcom Server) (include for Reference, Current and Temperature Shift Map data)

600d to 600e – data point location of end of cable/sensor – 0x xx Hex.

600f – 00 Hex = no ZBA, 01 to FF Hex = ZBA installed.

6010 to 6113 – Cable setup data – Setup data always begins and ends with 4 bytes of 00 Hex. (included in Cable/Sensor information on Palcom Server) (include for Reference, Current and Temperature Shift Map data)

- Each section of cable/sensor is packed into four bytes.

- High order nibble of 1st byte = cable/sensor type

- Low order nibble of 1st byte and 2nd byte = total counts from beginning to end of this section (include “total counts to end of all sections” for Reference, Current and Temperature Shift Map data)

- 3rd and 4th byte = total length (in meters or feet) from beginning to end of this section

- Section 1 is always located at address 6014 Hex

- Maximum number of section = 60

6114 to 6115 – not used

6116 to 6195 – Copy of original setup cable type and counts (data is saved only at the end of cable/sensor setup or at the end of recalibration).  Copy of setup data always begins and ends with 2 bytes of 00 Hex.

6196 – not used

6197 to 6198 – Force End (Should only be visible if enable and can only be editable by super admin with special passcode from PermAlert)

- MUST HAVE UPPER NIBBLE $A (TO FORCE IF NO SHORT OR BREAK FOUND) OR $E (TO FORCE ALWAYS), LOWER 3 NIBBLES COUNT TO FORCED AS END.

6199 to 619f – not used

61a0 – must be 00

61a1 to 61a2 – must be a copy of values in 600d to 600e – data point location of end of cable/sensor – 0x xx Hex.

61a3 to 61b7– not used

61b8 to 61f6 – Optional output relay module table (included in Cable/Sensor information on Palcom Server) –

- Memory location is the relay number.

- The contents of the memory location contain the section of cable that the relay is assigned to

- Example – if memory location 61b8 contained 05 Hex and an alarm was detected in section 5 relay 1’s contacts would switch states.

- If the content = 00 Hex then the relay is not assigned.

- If the content = valid section data (01 to 3c Hex), then include that in the Cable/Sensor setup information.

All the Cable setup and flag data can be retrieved as one string or they can be retrieved individually.

See examples below:

To retrieve all Cable 3 Setup and Flag Data, enter – DP,3,1,6000,61ff

MON 30A>

1/18/2019 08:50:13.099 [TX] - DP,3,1,6000,61ff

1/18/2019 08:50:13.125 [RX] - DP,3,1,6000,61ff<CR><LF>

6000  02 00 08 00  7D FF 11 20  FF AA 55 04  00 00 69 00  ....}.. .*U...i.<CR><LF>

6010  00 00 00 00  30 0C 00 37  D0 18 00 7E  30 2E 00 E3  ....0..7P..~0..c<CR><LF>

6020  D0 51 01 B2  A0 5E 01 ED  30 69 02 1F  00 00 00 00  PQ.2 ^.m0i......<CR><LF>

6030  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6040  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6050  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6060  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6070  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6080  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6090  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

60A0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

60B0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

60C0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

60D0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

60E0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

60F0  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6100  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6110  FF FF FF FF  FF FF 00 00  30 0C D0 18  30 2E D0 51  ........0.P.0.PQ<CR><LF>

6120  A0 5E 30 69  00 00 FF FF  FF FF FF FF  FF FF FF FF   ^0i............<CR><LF>

6130  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6140  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6150  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6160  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6170  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6180  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6190  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

61A0  00 00 69 FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ..i.............<CR><LF>

61B0  FF FF FF FF  FF FF FF FF  00 00 00 00  00 00 00 00  ................<CR><LF>

61C0  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

61D0  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

61E0  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

61F0  00 00 00 00  00 00 00 00  FF FF FF FF  FF FF FF FF  ................<CR><LF>

MON 30A>

The above setup and flag data shows that Cable 3 has 6 sections, is set to Medium sensitivity, is using criteria table 4, is monitoring against Map 1, has monitoring enabled, is monitoring without a ZPA installed, has no optional output relays assigned.

Force end is not enabled (keep hidden).

To retrieve Cable 3 Sensitivity Data, enter – DP,3,1,6006,6006

MON 30A>

1/18/2019 08:53:18.488 [TX] - DP,3,1,6006,6006

1/18/2019 08:53:18.517 [RX] - DP,3,1,6006,6006<CR><LF>

6006  11  .<CR><LF>

MON 30A>

Cable 3 Medium = 11 Hex

To retrieve Cable 3’s last data point that is recorded, enter – DP,3,1,6003,6004

MON 30A>

1/18/2019 08:54:33.108 [TX] - DP,3,1,6003,6004

1/18/2019 08:54:33.156 [RX] - DP,3,1,6003,6004<CR><LF>

6003  00 7D  .}<CR><LF>

MON 30A>

Cable 3 last data point = 007d Hex

To retrieve Cable 3 Section 1 Setup Data, enter – DP,3,1,6014,6017

MON 30A>

1/18/2019 09:01:15.348 [TX] - DP,3,1,6014,6017

1/18/2019 09:01:15.368 [RX] - DP,3,1,6014,6017<CR><LF>

6014  30 0C 00 37  0..7<CR><LF>

MON 30A>

Cable 3 Section 1 setup data decoded –

- Cable type for section 1 = 3 hex (3 decimal) (from high order nibble of 1st byte)

- Counts from panel to end of section 1 = 00C hex (12 decimal) from Low order nibble of 1st byte and 2nd byte

- Length from panel to end of section 1 = 0037 hex (55 decimal) from 3rd and 4th byte

To retrieve Cable 3 Section 4 Setup Data, enter – DP,3,1,6020,6023

MON 30A>

1/18/2019 09:03:56.087 [TX] - DP,3,1,6020,6023

1/18/2019 09:03:56.143 [RX] - DP,3,1,6020,6023<CR><LF>

6020  D0 51 01 B2  PQ.2<CR><LF>

MON 30A>

Cable 3 Section 4 setup data decoded –

- Cable type for section 4 = d hex (13 decimal) (from high order nibble of 1st byte)

- Counts from panel to end of section 4 = 051 hex (81 decimal) from Low order nibble of 1st byte and 2nd byte

- Length from panel to end of section 4 = 01b2 hex (434 decimal) from 3rd and 4th byte

The setup and flag data for the above cable 3 example generated the following Palcom 10 screen shot:

When retrieving Cable Setup and Flag data is completed enter AD,ff (this will put the panel into listen only mode)

END Cable x Setup and Flag Data (x= 1 to 4)

The following is required to create cable map graphs:

Retrieve Cable setup and flag data as shown above.

To retrieve Cable 3 Current Map data, enter dp,3,1,4040,40bd (40bd = 4040 + Cable 3, last data point 007d)

MON 30A>

1/18/2019 09:53:36.368 [TX] - dp,3,1,4040,40bd

1/18/2019 09:53:36.394 [RX] - dp,3,1,4040,40bd<CR><LF>

4040  00 00 00 00  00 00 00 00  22 20 23 22  21 21 20 1E  ........" #"!! . <CR><LF>

4050  1F 1F 1F 1F  22 21 23 21  22 23 22 21  22 21 20 1F  ...."!#!"#"!"! .<CR><LF>

4060  20 20 20 20  20 20 20 20  20 20 20 20  20 20 22 27                "'<CR><LF>

4070  2A 2B 2B 2C  2C 28 24 22  21 21 21 20  20 20 20 20  *++,,($"!!!     <CR><LF>

4080  20 20 20 20  20 20 20 20  20 20 20 20  21 21 20 20              !!  <CR><LF>

4090  20 1F 1D 1D  19 14 0F 08  04 03 04 03  04 0B 10 14   ...............<CR><LF>

40A0  19 1B 1C 1E  1F 20 20 20  20 24 3F 3F  3F 3F 3F 3F  .....    $??????<CR><LF>

40B0  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F 3F 3F  3E 39  ????????????>9<CR><LF>

MON 30A>

To retrieve Cable 3 Reference Map 1 data, enter dp,3,1,6218,6295 (6295 = 6218- + Cable 3, last data point 007d)

MON 30A>

1/18/2019 09:50:46.794 [TX] - dp,3,1,6218,6295

1/18/2019 09:50:46.829 [RX] - dp,3,1,6218,6295<CR><LF>

6218  AA 13 01 12  FF FF FF FF  21 20 23 22  21 21 20 1E  *.......! #"!! .<CR><LF>

6228  1F 1F 1F 1F  22 21 23 21  22 23 21 21  22 21 20 1F  ...."!#!"#!!"! .<CR><LF>

6238  20 20 20 20  20 20 20 20  20 20 20 20  20 20 22 27                "'<CR><LF>

6248  2A 2B 2B 2C  2C 28 24 22  21 21 21 20  20 20 20 20  *++,,($"!!!     <CR><LF>

6258  20 20 20 20  20 20 20 20  20 20 20 20  21 20 20 20              !   <CR><LF>

6268  20 1F 1D 1D  1C 1C 1C 1C  1C 1E 1F 20  20 20 20 20   ..........     <CR><LF>

6278  21 22 22 23  23 24 24 24  23 28 3F 3F  3F 3F 3F 3F  !""##$$$#(??????<CR><LF>

6288  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F  ??????????????<CR><LF>

MON 30A>

To retrieve Cable 3 Reference Map 2 data, enter dp,3,2,6218,6295 (6295 = 6218- + Cable 3, last data point 007d)

MON 30A>

1/18/2019 09:57:51.848 [TX] - dp,3,2,6218,6295

1/18/2019 09:57:51.907 [RX] - dp,3,2,6218,6295<CR><LF>

6218  CC 13 01 12  FF FF FF FF  22 21 22 22  21 21 20 1E  L......."!""!! .<CR><LF>

6228  1F 20 1F 1F  22 21 23 21  22 23 22 21  22 21 20 1F  . .."!#!"#"!"! .<CR><LF>

6238  20 20 20 20  20 20 20 20  20 20 20 20  20 20 22 27                "'<CR><LF>

6248  2A 2B 2B 2C  2C 28 24 22  21 21 21 20  20 20 20 20  *++,,($"!!!     <CR><LF>

6258  20 20 20 20  20 20 20 20  20 20 20 20  21 21 20 20              !!  <CR><LF>

6268  20 1F 1D 1D  19 14 0F 08  04 03 04 03  04 0A 10 14   ...............<CR><LF>

6278  19 1B 1D 1E  1F 20 20 20  20 24 3F 3F  3F 3F 3F 3F  .....    $??????<CR><LF>

6288  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F 3F 3F  3E 39  ????????????>9<CR><LF>

MON 30A>

To retrieve Cable 3 Reference Map 3 data, enter dp,3,3,6218,6295 (6295 = 6218- + Cable 3, last data point 007d)

MON 30A>

1/18/2019 10:11:09.596 [TX] - dp,3,3,6218,6295

1/18/2019 10:11:09.625 [RX] - dp,3,3,6218,6295<CR><LF>

6218  DD 12 0B 1D  FF FF FF FF  21 20 23 22  21 21 21 1F  ].......! #"!!!.<CR><LF>

6228  20 20 1F 1F  22 21 23 21  23 23 22 21  23 22 20 1F    .."!#!##"!#" .<CR><LF>

6238  20 20 20 20  20 20 20 20  20 20 20 20  20 20 22 27                "'<CR><LF>

6248  29 2A 2A 2A  2A 28 23 22  21 21 21 20  20 20 20 20  )****(#"!!!     <CR><LF>

6258  20 20 20 20  20 20 20 20  20 20 20 20  21 21 20 20              !!  <CR><LF>

6268  20 1F 1E 1D  1D 1C 1D 1C  1D 1E 1F 1F  20 20 20 21   ...........   !<CR><LF>

6278  23 26 27 28  29 2A 2A 29  26 29 3F 3F  3F 3F 3F 3F  #&'()**)&)??????<CR><LF>

6288  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F  ??????????????<CR><LF>

MON 30A>

To retrieve Cable 3 Reference Map 4 data, enter dp,3,4,6218,6295 (6295 = 6218- + Cable 3, last data point 007d)

MON 30A>

1/18/2019 10:21:56.626 [TX] - dp,3,4,6218,6295

1/18/2019 10:21:56.699 [RX] - dp,3,4,6218,6295<CR><LF>

6218  CC 12 0B 1D  FF FF FF FF  22 20 23 22  21 21 21 1F  L......." #"!!!.<CR><LF>

6228  20 20 1F 1F  22 21 23 21  23 23 22 21  23 21 20 1F    .."!#!##"!#! .<CR><LF>

6238  20 20 20 20  20 20 20 20  20 20 20 20  20 20 22 27                "'<CR><LF>

6248  29 2A 2A 2A  2A 28 23 22  21 21 21 20  20 20 20 20  )****(#"!!!     <CR><LF>

6258  20 20 20 20  20 20 20 20  20 20 20 20  21 20 20 20              !   <CR><LF>

6268  20 1F 1D 1D  1C 1C 1C 1C  1C 1E 1F 1F  20 20 20 20   ...........    <CR><LF>

6278  22 23 24 25  25 25 25 25  24 28 3F 3F  3F 3F 3F 3F  "#$%%%%%$(??????<CR><LF>

6288  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F  ??????????????<CR><LF>

MON 30A>

To retrieve Cable 3 Reference Map 5 data, enter dp,3,4,6218,6295 (6295 = 6218- + Cable 3, last data point 007d)

MON 30A>

1/18/2019 10:22:30.985 [TX] - dp,3,5,6218,6295

1/18/2019 10:22:31.030 [RX] - dp,3,5,6218,6295<CR><LF>

6218  CC 12 0B 1D  FF FF FF FF  22 20 23 22  21 21 21 1F  L......." #"!!!.<CR><LF>

6228  20 20 1F 1F  22 21 23 21  23 23 22 21  23 22 20 1F    .."!#!##"!#" .<CR><LF>

6238  20 20 20 20  20 20 20 20  20 20 20 20  20 20 22 27                "'<CR><LF>

6248  29 2A 2A 2A  2B 28 23 22  21 21 21 20  20 20 20 20  )***+(#"!!!     <CR><LF>

6258  20 20 20 20  20 20 20 20  21 20 20 20  21 21 20 20          !   !!  <CR><LF>

6268  20 1F 1D 1D  19 14 0F 09  05 03 04 03  04 0A 10 14   ...............<CR><LF>

6278  19 1B 1D 1E  1F 20 20 21  20 24 3F 3F  3F 3F 3F 3F  .....  ! $??????<CR><LF>

6288  3F 3F 3F 3F  3F 3F 3F 3F  3F 3F 3F 3F  3F 3A  ?????????????:<CR><LF>

MON 30A>

To retrieve Cable 3 Reference Map 6 data, enter dp,3,4,6218,6295 (6295 = 6218- + Cable 3, last data point 007d)

If data is all FF Hex the Reference Map was not used and does not need to be displayed.

MON 30A>

1/18/2019 10:22:53.433 [TX] - dp,3,6,6218,6295

1/18/2019 10:22:53.455 [RX] - dp,3,6,6218,6295<CR><LF>

6218  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6228  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6238  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6248  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6258  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6268  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6278  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6288  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF  ..............<CR><LF>

MON 30A>

To retrieve Cable 3 Reference Map 7 data, enter dp,3,4,6218,6295 (6295 = 6218- + Cable 3, last data point 007d)

If data is all FF Hex the Reference Map was not used and does not need to be displayed.

MON 30A>

1/18/2019 10:23:19.169 [TX] - dp,3,7,6218,6295

1/18/2019 10:23:19.195 [RX] - dp,3,7,6218,6295<CR><LF>

6218  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6228  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6238  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6248  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6258  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6268  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6278  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF FF FF  ................<CR><LF>

6288  FF FF FF FF  FF FF FF FF  FF FF FF FF  FF FF  ..............<CR><LF>

MON 30A>

To retrieve Cable 3 Temperature Shift Map data, enter dp,3,1,5040,50bd (50bd = 5040 + Cable 3, last data point 007d)

If data is all 00 Hex the Temperature Shift Map was not used and does not need to be displayed.

MON 30A>

1/18/2019 10:23:38.186 [TX] - dp,3,1,5040,50bd

1/18/2019 10:23:38.227 [RX] - dp,3,1,5040,50bd<CR><LF>

5040  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................ <CR><LF>

5050  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

5060  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

5070  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

5080  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

5090  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

50A0  00 00 00 00  00 00 00 00  00 00 00 00  00 00 00 00  ................<CR><LF>

50B0  00 00 00 00  00 00 00 00  00 00 00 00  00 00  ..............<CR><LF>

MON 30A>

For the following Palcom 10 Graph example please refer to the Cable x Setup and Flag Data memory location details.

Guidelines for creating graph from, from the Reference Map data, Current Map data, Temp Shift Map data, Cable x Setup and Flag Data.  Using above Cable 3 data to create the following graph.

- Save following table for setup table display information and for plotting x coordinates (feet or meters)

Sec.  Total  Total  Cable  Sect.  Sect.  Len/

No.   Len.   Cnts   Type   Len.   Cnts   Cnt

1     55     12      3     55     12   4.58

2    126     24     13     71     12   5.91

3    227     46      3    101     22   4.59

4    434     81     13    207     35   5.91

5    493     94     10     59     13   4.53

6    543    105      3     50     11   4.54

- First 8 data points on graph are not used or plotted.

- Calculate and save xy coordinates for each plotted Map:

- The total counts at the end of each section (12, 24, 46, 81, 94 and 105) must have an oversized plot point symbol (see included sample graph).

| Data point # (counts) | Length per count | Graph – X axis (feet) | Current Y axis | Map 1 Y axis | Map 2 Y axis | Map 3 Y axis | Map Y axis 4 | Map 5 Y axis | Map 6 Y axis | Map 7 Y axis | Temp Shift Y axis |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 - 8 | 4.59 | na | na | na | na | na | na | na | na | na | na |  |
| 9 | 4.59 | 41 | 22 | 21 | 22 | 21 | 22 | 22 | ff | ff | 00 |  |
| 10 | 4.59 | 46 | 20 | 20 | 21 | 21 | 20 | 20 | ff | ff | 00 |  |
| 11 | 4.59 | 50 | 23 | 23 | 22 | 23 | 22 | 22 | ff | ff | 00 |  |
| 12 | 4.59 | 55 | 21 | 21 | 21 | 21 | 21 | 21 | ff | ff | 00 |  |
| 13 | 5.93 | 61 | 21 | 21 | 21 | 21 | 21 | 21 | ff | ff | 00 |  |
| 14 | 5.93 | 67 | 20 | 20 | 20 | 21 | 21 | 21 | ff | ff | 00 |  |
| … | … | … | … | … | … | … | … | … | … | … | … |  |
| 44 | 4.59 | 218 | 20 | 20 | 20 | 20 | 20 | 20 | ff | ff | 00 |  |
| 45 | 4.59 | 222 | 20 | 20 | 20 | 20 | 20 | 20 | ff | ff | 00 |  |
| 46 | 4.59 | 228 | 22 | 22 | 22 | 22 | 22 | 22 | ff | ff | 00 |  |
| 47 | 5.93 | 235 | 27 | 27 | 27 | 27 | 27 | 27 | ff | ff | 00 |  |
| 48 | 5.93 | 240 | 2a | 2a | 2a | 29 | 29 | 29 | ff | ff | 00 |  |
| … | … | … | … | … | … | … | … | … | … | … | … |  |
| 80 | 5.93 | 429 | 20 | 20 | 20 | 20 | 20 | 20 | ff | ff | 00 |  |
| 81 | 5.93 | 434 | 1f | 1f | 1f | 1f | 1f | 1f | ff | ff | 00 |  |
| 82 | 4.59 | 439 | 1d | 1d | 1d | 1d | 1d | 1d | ff | ff | 00 |  |
| 83 | 4.59 | 443 | 1d | 1d | 1d | 1d | 1d | 1d | ff | ff | 00 |  |
| 84 | 4.59 | 448 | 19 | 1c | 19 | 1d | 1c | 19 | ff | ff | 00 |  |
| … | … | … | … | … | … | … | … | … | … | … | … |  |
| 104 | 4.59 | 538 | 20 | 23 | 20 | 26 | 24 | 20 | ff | ff | 00 |  |
| 105 | 4.59 | 543 | 24 | 28 | 24 | 29 | 28 | 24 | ff | ff | 00 |  |
| 106 | 4.59 | 548 | 3f | 3f | 3f | 3f | 3f | 3f | ff | ff | 00 |  |
| 107 | 4.59 | 553 | 3f | 3f | 3f | 3f | 3f | 3f | ff | ff | 00 |  |
| 108 | 4.59 | 558 | 3f | 3f | 3f | 3f | 3f | 3f | ff | ff | 00 |  |
| 109 | 4.59 | 563 | 3f | 3f | 3f | 3f | 3f | 3f | ff | ff | 00 |  |
| 110 | 4.59 | 168 | 3f | 3f | 3f | 3f | 3f | 3f | ff | ff | 00 |  |

Date information for map 1 through map 7 data (as shown on graph example) is located in the 2nd (last 2 digits of year in hex), 3rd (month in hex) and 4th (day in hex).  Current map date will be the day the graph data was downloaded from the PAL-AT panel.

In the following graph example, the graph data was gathered on 2019-01-18

Map 1 date memory locations, 6219-621B (13 01 12 hex) = 19-01-18 (displayed on graph as 2019-01-18)

Map 2 date memory locations, 6219-621B (13 01 12 hex) = 19-01-18 (displayed on graph as 2019-01-18)

Map 3 date memory locations, 6219-621B (12 0B 1D hex) = 18-11-29 (displayed on graph as 2018-11-29)

Map 4 date memory locations, 6219-621B (12 0B 1D hex) = 18-11-29 (displayed on graph as 2018-11-29)

Map 5 date memory locations, 6219-621B (12 0B 1D hex) = 18-11-29 (displayed on graph as 2018-11-29)

Map 6 date memory locations, 6219-621B (FF FF FF hex) = invalid – no map

Map 7 date memory locations, 6219-621B (FF FF FF hex) = invalid – no map

When retrieving Cable map data is completed enter AD,ff (this will put the panel into listen only mode)

END of create cable map graphs

Uploading Cable Setup data to PAL-AT:

To upload the cable 3’s setup data for system ID 004 (shown above) send the following ASCII commands:

Enter Esc key or 1B (hex code for the Esc key)

Terminal response if communication port is open:

1/21/2019 14:59:53.016 [TX] - <ESC>_

1/21/2019 14:59:53.021 [RX] - <NUL><CR><LF>

<NUL>MON 24M>

Terminal response if communication port is closed (listen only mode):

1/21/2019 15:00:30.487 [TX] - <ESC>_

To open communication to the AT30 series panel system ID 004, enter -  AD,04

Terminal response:

11/9/2018 13:52:30.187 [TX] - ad,04

11/9/2018 13:52:30.241 [RX] - <CR><LF>

MON 30A>

Enable restricted access, enter PW,3179636424

Terminal response:

11/8/2018 11:50:58.459 [TX] - pw,3179636424

11/8/2018 11:50:58.503 [RX] - pw,XXXXXXXXXX ENABLED<CR><LF>

MON 30A>

To restore cable/sensor data enter the following:

JF,3,600002 00 08 00 7D FF 11 1E 3F AA 55 04 01 00 69 00 00 00 00 00.

JF,3,601430 0C 00 37 D0 18 00 7E 30 2E 00 E3 D0 51 01 B2 A0 5E 01 ED 30 69 02 1F 00 00 00 00 .

JF,3,611600 00 30 0C D0 18 30 2E D0 51 A0 5E 30 69 00 00 .

JF,3,61A000 00 69.

To commit the cable/sensor data to non-volatile memory: enter jf,w

Terminal response:

MON 30A> JF,W<CR><LF>

MON 30A>

Panel must be reset: enter wd (note – after a successful wd communication to the panel must be reestablished)

Terminal response:

MON 30A> WD

The panel will take several seconds to reboot.

Verify that data was written to panel correctly:

To retrieve all Cable 3 Setup and Flag Data, enter – DP,3,1,6000,61ff

When uploading setup data is completed enter AD,ff (this will put the panel into listen only mode)

END of Uploading Cable Setup data to PAL-AT

Cable Status update data from PAL-AT:

To retrieve Cable status data, enter CS

Terminal response:

MON 30A>

1/24/2019 14:39:51.714 [TX] - cs

1/24/2019 14:39:51.745 [RX] - cs<CR><LF>

<CR><LF>

12000003<CR><LF>

21000003<CR><LF>

32000005<CR><LF>

42000006<CR><LF>

<CR><LF>

MON 30A>

CS Command

The “CS”, or Cable Status, command returns the status data for all cables.  When PAL-AT receives "CS", it responds with 4 lines of data.  Each line gives the status of each cable, the number of alarms in the alarm queue, and the distance to the problem (or probe number).  Each line is 8 bytes long plus a carriage return.  The 8 bytes in each line of the ASCII response is decoded as follows:

_ _ _ _ _ _ _ _+CR

1 2 3 4 5 6 7 8

Byte #  	Description

1	Cable # (1-4)

Byte #  	Description

2	Cable Status (defined below)

0=not installed

1=off-line

2=normal

3=leak

4=break

5=short

6=probe active

7=probe reset

8=drying

9=no end/fault in 50 ft [15 m]

Byte 	#  Description

3 	Number of alarms in alarm queue

4	Distance or probe number

5  	Distance or probe number

6

7

8  	(sum of bytes 1- 7) MOD 10

Note:  If the panel loses power when there are alarms in the alarm queue, the cable status will be “1” (off-line) after power is restored, until the alarm queue is cleared.  The reason for this is the system does not change the cable status to “2” (normal) while alarms are in the queue.  For example, when a leak is detected the cable status will be “3”.  It will remain “3”, even after the new map is created and the system resumes monitoring and displays “Cable OK”.  The system requires an authorized password to acknowledge the leak event occurred and then it displays the status.

END of Cable Status update

Not completed

Archive Memory Map and Archive Debug Commands

- Archive messages are packed into 8 bytes

- 928 Archives

- Stored in 4 - 800 byte blocks

- When the last archive in a block is written the next block will be erased. (if the last archive in block 4 is written block 1 will be erased)

Archive Pointer (displays pointer to next archive to be stored – will always be cleared (all FF’s))

To retrieve Archive Pointer, enter – AP

- 0000 = first archive memory location – 039f– last archive memory location

- When archive pointer is incremented past 039f it will reset to 0000

MON 30A>

12/4/2018 07:40:43.229 [TX] - ap

12/4/2018 07:40:43.253 [RX] - ap016D <CR><LF>

MON 30A>

The result of the example – 0166 (hex)

Archive Retrieval (AR,(hex number from 0000 to 039c)   WILL DUMP 4 ARCHIVES (32 CHARACTERS) AS STRING OF HEX)

To retrieve the last 4 archives, enter AR, Archive Pointer – 4 (using archive pointer from previous example).

MON 30A>

12/4/2018 07:44:03.083 [TX] - ar,0169

12/4/2018 07:44:03.106 [RX] - ar,01690C030E140C030E140C030F140C030F140C8310055F0232140C8310105F023214<CR><LF>

MON 30A>

To retrieve Archives, the first time (rough draft)

- Retrieve and store AP value in temp registers (TempAP) and (LastAP)

- Subtract 1 from TempAP (if result is less than 0000H set TempAP to 039FH)

- (Special Case) - If TempAP was equal to 039fH and the archive message pointed to by LastAP is all FF’s then there are no archive messages in memory (Give Archive Empty message)

- Retrieve last 4 archives (AR, TempAP)

MON 30A>

12/4/2018 07:42:29.555 [TX] - ar,016c

12/4/2018 07:42:29.600 [RX] - ar,016c0C8310105F023214FFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFF<CR><LF>

MON 30A>

- Save 1st 8 bytes (0C8310105F023214) and decode

- any of the 4 8 byte

- If any of the 4 8 byte packed archive messages are all FF’s then done gathering archives

- Subtract 4 from TempAP (if TempAP is 0000 or less set to 0380

- Stored in 4 - 800 byte blocks

- When the last archive in a block is written the next block will be erased. (if the last archive in block 4 is written block 1 will be erased)

;  1st Byte    |  2nd Byte    |  3rd Byte    |  4th Byte   |    5th Byte   |   6th Byte    |   7th Byte    |   8th Byte

;xxxx|XXXX  xxxX|XXXX  xxxX|XXXX  xxxx|XXXX  XXXX|XXXX XXXX|XXXX  XXXX|XXXX XXXX|XXXX

;|||| ||||    |||| ||||     |||| ||||     |||| ||||    ||||  ||||    ||||  ||||    ||||  ||||    ||||  ||||- year

;|||| ||||    |||| ||||     |||| ||||     |||| ||||    ||||  ||||    ||||  ||||    ||||  ||||- security level (tens and ones digits – not used when distance value in 5th and 6th bytes)

;|||| ||||    |||| ||||     |||| ||||     |||| ||||    ||||  ||||    ||||  ||||- Distance or security level (hundreds digits)

;|||| ||||    |||| ||||     |||| ||||     |||| ||||    ||||  ||||- Distance or 	# or Map# or Password#

;|||| ||||    |||| ||||     |||| ||||     |||| ||||- Minute Archive saved (Hex)

;|||| ||||    |||| ||||     |||| ||||     ||-

;|||| ||||    |||| ||||     |||| ||||- Hour Archive saved (Hex)

;|||| ||||    |||| ||||     |||- Secondary Subtype table (see detail of secondary subtype in 3rd byte detail below)

;|||| ||||    |||| ||||- Day Archive was saved(Hex)

;|||| ||||    |||-  Subtype table

;|||| ||||- Month Archive saved (Hex)

;||||- Cable Number or Type 0 (Type 0 is non cable related archive)

1st Byte - upper 4 bits = Archive type 0 or cable number (1 through 4)

1st Byte - lower 4 bits = month

2nd Byte – upper 3 bits = Subtype table

Use for Archive Type 0

Subtype 0 = Power failure

Subtype 1 = Time Set

Subtype 2 = Year Changed

Subtype 3 = Archives Cleared

Subtype 4 = Security Entry

Subtype 5 = Monitoring Enabled

Subtype 6 = Monitoring Disabled

Subtype 7 = use Secondary Subtype defined in 3rd Byte

Use for Cable 1 through 4 Archives (Not Type 0)

Subtype 0 = Drying

Subtype 1 = Leak

Subtype 2 = Probe xx Activated (xx = 1 through 31)

Subtype 3 = Probe xx Activated (xx = 32 through 63)

Subtype 4 = Probe xx Reset (xx = 1 through 31)

Subtype 5 = Probe xx Reset (xx = 32 through 63)

The Probe number for Subtype 2, 3, 4 and 5 is stored in the upper 3 bits of 3rd byte (low order bits) and the upper 2 bits of 	the 4th byte (high order bits)

Subtype 6 = use Secondary Subtype defined in 3rd Byte

Subtype 7 = use Secondary Subtype defined in 3rd Byte

2nd Byte – lower 5 bits = day of the month

3rd Byte -  upper 3 bits = Secondary Subtype table

Use for Archive Type 0 Subtype 7

Secondary Subtype 0 = Monitoring Time Expired

Secondary Subtype 1 = Unsuccessful Phone Call Reason

Secondary Subtype 2 = Successful Phone Call – Primary Number

Secondary Subtype 3 = Successful Phone Call – Secondary Number

Secondary Subtype 4 = Not Used

Secondary Subtype 5 = Not Used

Secondary Subtype 6 = Not Used

Secondary Subtype 7 = Not Used

Use for (Not Type 0) Subtype 2,3,4,5

Prove number bits 0-2

Use for (Not Type 0) Subtype 6

Secondary Subtype 0 = Fault in first 50 feet or Fault in first 15 meters

Secondary Subtype 1 = No end found in

Secondary Subtype 2 = Auto select map

Secondary Subtype 3 = Short detected

Secondary Subtype 4 = Auto-reference taken due to temperature shift

Secondary Subtype 5 = Not Used

Secondary Subtype 6 = Not Used

Secondary Subtype 7 = Not Used

Use for (Not Type 0) Subtype 7

Secondary Subtype 0 = Monitoring on

Secondary Subtype 1 = Monitoring off

Secondary Subtype 2 = New setup taken

Secondary Subtype 3 = New reference taken

Secondary Subtype 4 = Auto-reference taken due to drying

Secondary Subtype 5 = Acknowledged

Secondary Subtype 6 = Break detected

Secondary Subtype 7 = Normal

3rd Byte – Lower 5 bits = hour

4th Byte – upper 2 bits = Probe Bits 3,4

4th Byte – lower 6 bits = minutes

5th and 6th byte – distance

Or

5th byte – Password number from table – can be 1 to 60 or 95 for Demo Password

Or

6th byte – Map# or security level (hundreds digits)

7th byte – security level (ten’s and one’s digits)

8th byte – Year (ten’s and one’s digits)

MAP NUMBER KEPT AS 0-4, SHOW 1-5 IN ARCHIVE

; ARCHIVE TYPE STORED IN 4 BITS ABOVE MONTH (IN FIRST ARCHIVE BYTE)

; ARCHIVE SUBTYPE STORED IN 3 BITS ABOVE DAY (IN SECOND ARCHIVE BYTE)

; TYPE ZERO SUBTYPE TABLE:

ZROTYP:     DC.W  HISTX5      ; SUBTYPE 0 - POWER FAILURE

DC.W  HISTX6      ; SUBTYPE 1 - TIME SET

DC.W  HISTX8      ; SUBTYPE 2 - YEAR CHANGED

DC.W  HISTX9      ; SUBTYPE 3 - ARCHIVES CLEARED

DC.W  HISTXF      ; SUBTYPE 4 - SECURITY ENTRY

DC.W  HISCT5      ; SUBTYPE 5 - MONITORING ENABLED

DC.W  HISCT6      ; SUBTYPE 6 - MONITORING DISABLED

; SUBTYPE 7 BROKEN INTO 8 SECONDARY SUBTYPES

;  SECONDARY SUBTYPE STORED IN 3 BITS ABOVE HOUR (IN THIRD ARCHIVE BYTE)

ZROTYPend:  equ   $

BASE07:     EQU   ZROTYPend-ZROTYP    ; BASE FOR SUBTYPE 7, SECONDARY SUBTYPE 0

DC.W  HISCT7      ; SECONDARY SUBTYPE 0 - MONITORING TIME LIMIT EXPIRED

DC.W  HISPH1      ; SECONDARY SUBTYPE 1 - UNSUCCESSFUL PHONE CALL REASON

DC.W  HISPH2      ; SECONDARY SUBTYPE 2 - SUCCESSFUL PHONE CALL-PRIMARY NUMBER

DC.W  HISPH3      ; SECONDARY SUBTYPE 3 - SUCCESSFUL PHONE CALL-ALTERNATE NUMBER

DC.W  HISTX4      ; INVALID SUBTYPE (SECONDARY SUBTYPE 4)

DC.W  HISTX4      ; INVALID SUBTYPE (SECONDARY SUBTYPE 5)

DC.W  HISTX4      ; INVALID SUBTYPE (SECONDARY SUBTYPE 6)

DC.W  HISTX4      ; INVALID SUBTYPE (SECONDARY SUBTYPE 7)

; TYPES 1-8 (CABLE NUMBER) SUBTYPE TABLE:

; IF NOT SUBTYPES 6 OR 7 5 BIT DATA CAN BE STORED ABOVE BYTES 3 AND 4

;  LOW 3 BITS ABOVE HOUR (THIRD BYTE) AND NEXT 2 BITS ABOVE MINUTE (FOURTH BYTE)

;   FOR SUBTYPES 2 THRU 5 ADDITIONAL DATA IS PROBE NUMBER - ALLOWS PROBES 0-63

HISTYP:     DC.W  TXTVE7      ; SAY DRYING (SUBTYPE 0)

DC.W  TXTVE8      ; SAY LEAK (SUBTYPE 1)

DC.W  TXTVEP      ; SAY PROBE 1-31 ACTIVATED (SUBTYPE 2)

DC.W  TXTVEP      ; SAY PROBE 32-63 ACTIVATED (SUBTYPE 3)

DC.W  TXTVER      ; SAY PROBE 1-31 RESET (SUBTYPE 4)

DC.W  TXTVER      ; SAY PROBE 32-63 RESET (SUBTYPE 5)

; SUBTYPE 6 BROKEN INTO 8 SECONDARY SUBTYPES

;  SECONDARY SUBTYPE STORED IN 3 BITS ABOVE HOUR (IN THIRD ARCHIVE BYTE)

HISTYPend1: equ   $

BASE60:     EQU   HISTYPend1-HISTYP    ; BASE FOR SUBTYPE 6, SECONDARY SUBTYPE 0

DC.W  TXTVEF      ; SAY FAULT IN FIRST 50 FEET (SECONDARY SUBTYPE 0)

DC.W  TXTSUZ      ; SAY NO END FOUND IN X FEET (SECONDARY SUBTYPE 1)

DC.W  HISTXS      ; SAY AUTO SELECT MAP (SECONDARY SUBTYPE 2)

DC.W  TXTVEY      ; SAY SHORT DETECTED (SECONDARY SUBTYPE 3)

DC.W  HISTXT      ; SAY AUTOREFERENCE TAKEN DUE TO TEMP SHIFT (SECONDARY SUBTYPE 4)

DC.W  HISTX4      ; INVALID SUBTYPE (SECONDARY SUBTYPE 5)

DC.W  HISTX4      ; INVALID SUBTYPE (SECONDARY SUBTYPE 6)

DC.W  HISTX4      ; INVALID SUBTYPE (SECONDARY SUBTYPE 7)

; SUBTYPE 7 BROKEN INTO 8 SECONDARY SUBTYPES

;  SECONDARY SUBTYPE STORED IN 3 BITS ABOVE HOUR (IN THIRD ARCHIVE BYTE)

HISTYPend2: equ   $

BASE70:     EQU   HISTYPend2-HISTYP    ; BASE FOR SUBTYPE 7, SECONDARY SUBTYPE 0

DC.W  HISTXC      ; SAY MONITORING ON (SECONDARY SUBTYPE 0)

DC.W  HISTXD      ; SAY MONITORING OFF (SECONDARY SUBTYPE 1)

DC.W  HISTXE      ; SAY NEW SETUP TAKEN (SECONDARY SUBTYPE 2)

DC.W  HISTXB      ; SAY NEW REFERENCE TAKEN (SECONDARY SUBTYPE 3)

DC.W  HISTXM      ; SAY AUTOREFERENCE TAKEN DUE TO DRYING (SECONDARY SUBTYPE 4)

DC.W  HISTX7      ; SAY ACKNOWLEDGED (SECONDARY SUBTYPE 5)

DC.W  TXTVES      ; SAY BREAK DETECTED (SECONDARY SUBTYPE 6)

DC.W  HISTX3      ; SAY NORMAL (SECONDARY SUBTYPE 7)

Legacy LiquidWatch ASCII protocol Debug Commands index:

; COMMAND TABLE

ZDATA6	EQU	$

DW	'DD'	;DUMP DATA

DW	ZDUDAT

DW	'DM'	;DUMP MEMORY

DW	ZDUMEM

DW	'PW'

DW	ZPWCK

DW	'AD'	;ADDRESS CHECK

DW	ZAD

DW	'DI'	;DISPLAY REQUEST

DW	ZDISP

DW	'KP'	;KEYPAD INPUT

DW	ZKEYP

DW	'AP'	;DUMP ARCHIVE POINTER

DW	ZARPTR

DW	'RA'	;REPLY ADDRESS (AS SET BY SA,XX)

DW	ZREPAD

DW	'AR'	;DUMP ARCHIVES

DW	ZDUMAR

DW	'PV'	;PROGRAM VERSION

DW	ZDUVER

DW	'PS'	;Dumps the status of all 64 Probes

DW	ProbeStat

ZSEC	DW	'CM'	;CLEAR MEMORY

DW	ZCLMEM

DW	'FM'	;FILL MEMORY

DW	ZFLMEM

DW	'IC'	;INSPECT/CHANGE

DW	ZINSCH

DW	'JM'	;JAM MEMORY (DOESN'T READ IT)

DW	ZJMMEM

DW	'CF'	;CLEAR FLAG

DW	ZCLRFL

;PS Command

;Dumps the status of 64 probes

;XX(Probe number 01-64)X(Probe Statis: 0-Active, 1 - Break, 2 - Short, 3 - Good,

;4 - Offline, 5 - Not Setup)XXX(Probe reading diagnostic number)X(MON10 Check Sum,

;{Decimal Addition of first six digits, send out least significat digit as seventh

;number})

Legacy LiquidWatch Archive decoding:

;Display archived Information

;Coded Archive message:

;       1st                 2nd                3rd                4th                 5th

;XXxX|XXXX  xxXX|XXXX  xxxX|XXXX  xxxx|XXXX  xXXX|XXXX

;|||| ||||  |||| ||||  |||| ||||  |||| ||||  |||| ||||

;|||| ||||  |||| ||||  |||| ||||  |||| ||||  |Year & Decade Archive was saved(Hex)

;|||| ||||  |||| ||||  |||| ||||  |||| ||||  Part of Centery Code

;|||| ||||  |||| ||||  |||| ||||  |||| Month Archive was saved(Hex)

;|||| ||||  |||| ||||  |||| ||||  Low Order Nibble of Probe Number(Hex)

;|||| ||||  |||| ||||  |||Day Archive was saved(Hex)

;|||| ||||  |||| ||||  High Order Nibble of Probe Number

;|||| ||||  ||Minute Archive was saved(Hex)

;|||| ||||  Fault Type

;|||Hours Archive was saved(Hex)

;||Part of Centery Code

;Archive Type

;

;Centery Code 00=1900's 80=2000's 20=2100's A0=2200's or greater

;Fault Type C0 or 40=Probe Break Message 80=Probe Shorted Message

;           00=Probe Active Message

;Archive Type C0=Power Lost/Restored/End of Archive Message 40=Probe Off Message

;             80=Probe On Message 00=Probe Alarm Message

;Probe Number 1 thru 64=Probe Number 7D=Power On Archive Message

;      7E=Power Off Archive Message (All other codes=End of Archive message)